영 문 규 격 서

Commodity & Description

관세분류번호 HSK NO.	품목번호 ITEM NO.	정부물품분류번호(8자리) Korean Government Commodity Classification Code(eight-digit)	품목 및 규격 DESCRIPTION	수량 Q'TY	단위 UNIT
9027.30.1000	1	41115408	FT-IR Microscope Imaging Spectroscopy	1	sys

A. End-user's Use

1. This system is a powerful tool for analyzing the sample of organic compounds, organometal,

polymer, wafer, ceramics, thin film, coatings, fiber, medicine, medical supplied, agricultural

products as well as unknown compound.

2. This system should be sealed and desiccated but purge system should be standard

3. This system can measure the sample size as small as 1.56um

4. Transmission, reflectance, ATR and Imaging ATR modes

5. Single-point, multi-point, linescan, mapping, ATR mapping and autofocus mapping

6. The manufacturer must have ISO 9001 certification for the design, manufacture, and service of the

FT-IR instrument.

7. A unique and powerful microscopy system optimized for both IR and visual performance, with

all microscope components multimedia software. Include motorized sample stage, integral CCD video

camera and video capture card for the PC.

B. Configuration of Goods

1. FT-IR Spectrometer	1 set

2. Fully Automated Imaging system                                                      1 set

3. Accessories	1 set

4. Data Control System                                                                   1 set

C. Performance and Specification

1.  FT-IR Spectrometer                                                                   1 set

1) Standard Wavelength range ; Wider than Full IR Range 14,700 ～ 30cm-1

2) Interferometer : Improved Michelson interferometer Dynascan-Dynamic Alignment

3) Source : Pre-aligned tungsten halogen lamp for Near IR

Pre-aligned Mid, Far IR long life souce with exclusive electronic hot spot stabilization.

4) Detector : Temperature Stabilized Fast Recovery DTGS or InGaAs optimized for Near IR

Temperature Stabilized Fast Recovery DTGS optimized for Mid IR

Temperature Stabilized Fast Recovery DTGS or Polyethylene optimized for Far IR

5) Beamsplitter : Wide-range multi-layer calcium fluoride beamsplitter for the Near IR spectral region

Multi-layer potassium bromide beamsplitter for the Mid IR spectral region

Aluminum Grid / polypropylene beamsplitteer for the Far IR spectral resion

6) Automated Range Switching (ARS) for single button range change incorporating automatic change

of source, beamsplitter and detector

7) General Purpose Optical Bench with Photoacustic PA301 and external beam mirror

8) Resolution : Less than 1cm-1 at 10,000 cm-1 / 0.4cm-1 at 3,028 cm-1 band in Methane

9) Signal to Noise : Better than 174,800:1 RMS, 50,000/1 p-p

10) Wavelength Precision ; Better than 0.1cm-1 at 6,000cm-1 / ±0.008cm-1 at 2,000 cm-1

11) Real time Atmospheric correction.

12) Sealed and desiccated optics

13) Automatically sensing ATR. DR. etc.

14) Instrument Performance Verification(IPV)

15) Absolute Virtual Instrument (AVI)

16) Automatic Vapor Compensation (AVC)

2. Fully Automated Imaging system                                                        1 set

1) Alignment : Automatic Interferometer Management System full alignment system for

all instrument part.

2) Two in one operation : FT-IR imaging system and microscope built into the same system.

3) Detector : Duet detector containing Array MCT detector and MB MCT detector in same dewar,

on the same substrate. Mode selected with a single mouse click, with no mechanical

change over or re-alignment necessary.

Integrated, high performance stabilized InGaAs detector for Near IR

4) Operating modes :

- Transmission

- Reflectance

- Micro-ATR mode(Single point ATR).

- 1,100㎛ contact Imaging ATR mode(Imaging ATR).

5) Array operation : High performance linear MCT array with infinitely variable image size.

6) Scan speed : collect 170 high-quality spectra per second

7) Imaging Pixel size :

- 1.56um (Imaging ATR mode), 6.25㎛, 25㎛, 50㎛

8) Number of Dead Pixels : Zero

9) Spectral range :

- Single point frequency range : 10,000 ～ 600cm-1

- Array frequency range : 7,800 ～ 680cm-1

- 1,100㎛ ATR Imaging frequency range : 7,800 ～ 680cm-1

10) Sensitivity : 50,000/1(peak to peak), at 4cm-1 ,2min measurement.

11) Visible Illumination : Solid-state white light LED illumination provides more even

and accurate color, better brightness and longer life.

12) System must provide full mouse-controlled automation :

- Auto-focus, Auto-correction, Auto-Illumination.

- Auto-Transmission/Reflectance switching.

- Auto-aperturing.

13) Removable lower cassegrain objective to allow the measurement of thick samples.

14) Must be able to see IR image on-screen during the image capture process.

15) Must have multiple IR image and line-scan display modes :

- Interpolated false color.

- false color/contour,

- 3-D wire grid and 3-D rendered surface

16) Motorized sample stage : 3x2-inch(75x50mm) full travel. 1μm minimum step size.

17) Auto Aperture or ROHS Calibration curve Motorized aperture blades automatically conform to

size and shape of a 'growbox' defined directly on the visible image.

18) Auto Correction : Automatic adjustment of motorized lower cassegrain to optimize IR energy

for transmission measurement

19) Interactive Multimedia Software : Integral video camera provides on-screen visible image.

Sampling points for all data collection modes defined by

direct mouse interaction with live visible image

20) Joystick : Additional stage control, including focus, via proportional control joystick

21) Automated Data Collection : Automated measurement of individual point,

groups of point, line scan and maps. All point, lines and

map areas can be defined interactively on the video image

22) Data Traceability : Video image can be with stage coordinates and numbered

sample and background makers. Spectra are saved with coordinates

and maker numbers.

23) IR and visible beam use identical beam paths for optimum geometric image accuracy.

24) Integral video camera provides on-screen visible image.  Sampling points for all data collection

modes are defined by direct interaction with live visible image.

25) Key microscope and instrument components are checked on startup. Status of all components is

displayed via PC control toolbox.

26) Visible Polarizers

27) IR Polarizers.

28) Software

- Instrument Control and Data Manipulation Software

- Spectrum searching Software

- Spectrum Image Software

- 230,000ea Library

29) Micro Accessory Kit

- Micro Sampling Kit

- Gold Mirror Assemble

- Sample Holder (3x13mm)

- KBr Windows (2ea)

- Microscope Slide (72)

- Rotable Cell Holder

3. Accessories

1) Single Bounce ATR for Main system                                                   1 set

- Diamond/KRS5 (Cutoff : 380cm-1) crystal

- Automatic recognition

- Automatic optimization

- Zero alignment, zero set-up

- Pressure arm with pressure sensor(gauge) supplied as standard

2) Ge crystal for Single Bounce ATR                                                      1 ea

3) Imaging ATR system accessory                                                        1 set

- High Quality Optical and Mechanical Assemblies

- Aligned Ge Crystal

- Measurement Area : 1,1100㎛ at one contact

- Pixel resolution : 1.56㎛ (High Definition Mode)

4) Spare imaging ATR crystal (1,100x1,100㎛ Area) (Ge Crystal)                            1 ea

5) Micro ATR Objective for IR Micro system                                              1 ea

- Mount

- ATR Crystal Holder

- Ge ATR Crystal

6) Leica DMS 1000 Digital Microscope                                                     1 set

- zoom ratio : 8:1

- magnification range : 18x – 150x with 22"-monitor

- optical resolution : up to 159 lp/mm

- light intensity controlled LED ring

- microscope must have a parfocal optical system

- microscope must have a internal digital camera

- internal digital camera resolution : up to 2592×1944 (Aptina 1/2.3" CMOS)

- internal digital camera can be connect the monitor directly

- internal digital camera observe image, save image and save video without PC

- internal digital camera must have HDMI port

- internal digital camera must have SD card slot

- internal digital camera can be controlled with remote controller.

- internal digital camera speed : up to 30frame / 1sec

- Electronic Interface : USB 2.0, standard USB-plug type B

- fully coded zoom for measurement stand alone or with PC

- column must have the travel range up to 300 mm under control by coarse/fine focus drive

- depth of focus range : 3.5mm – 0.05mm

- microscope must have a motorized scanning stage

- motorized scanning stage is moving up to 15.24 x 10.16cm

- motorized scanning stage have a accuracy +/- 2 um

7) DMS 1000 System Software w/ operating workstation(PC)                                1 set

- measurement and annotation function

- software have a report function compatible with Exel

- measurement sample dimension without calibration

- multiple tile images compose to one bigger image (controlled by scanning stage)

- automated sophisticated measurements on individual features including size, shape,

position, orientation and intensity

4. Data Control System                                                                      1 set

- Microprocessor : Intel®i7 3.4 GHz Quad Core Processer

- RAM memory : 8GB or better

- Hard disk capacity : 1TB or better

- 22" LED monitor

- Windows 7 Pro 64-bit

- Laser Printer

D. Remarks

1. Two years warranty service should be provided supplier after the installation.

2. The education of maintenance and operation should be carries out by bidder/supplier's responsibility at end user's site.

3. Bidder/Supplier supply manual CD and paper written Korean and English.